## QC Test Summary-iCRco
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – iCRco
Site Name | [7 empty] | Report Date
Address | [7 empty] | Survey Date
Medical Physicist's Name | [7 empty] | Signature
X-Ray Unit Manufacturer | [7 empty] | Model
CR Image Receptor Mfr |  | iCRco | [5 empty] | CR Model
Date of Installation | [7 empty] | Room ID
QC Manual Version # | [6 empty] | (use any version applicable to model; contact mfr if questions)
Accessory Equipment |  | Manufacturer |  |  | Model |  | Location |  | QC Manual Version #
Review Workstation*
Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System www.fda.gov/CDRH/MAMMOGRAPHY/robohelp/START.HTM.
Survey Type |  | Mammo Eqpt Evaluation of new unit (include MQSA Rqmts for Mammo Eqpt checklist) | [9 empty] | Annual Survey
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
PASS/FAIL
1. AEC Calibration
2. Contrast-To-Noise Ratio Measurements and Signal-To-Noise Ratio (values required for all tests)
SNR (value)
| CNR (value) |  | (Required for both new unit Mammography Equipment Evaluations and Annual Surveys)
CNR should not vary by more than ±15% (NA for Equipment Evaluation)
3. Evaluation of System Resolution
4. Phantom Image Quality Evaluation
| Phantom image scores: |  |  | Fibers |  | Specks |  | Masses
5. Breast Entrance Exposure
6. Average Glandular Dose
| Average glandular dose for average breast is ≤3 mGy (300 mrad) | [8 empty] | mrad
7. Mammographic Unit Assembly Evaluation
8. Collimation Assessment
9. Beam Quality Assessment - HVL Measurement
10. AEC Function Performance/Tracking Reproducibility
11. kVp Accuracy and Reproducibility
12. Scanner Performance Test and Jitter
13. Inter-Plate Consistency
14. Artifact Evaluation
15. Modular Transfer Function (MTF) |  |  |  |  | (Tests 15, 16 and 17 are only required "when using the alternate test procedure for the Dynamic Range test and is included in the Dynamic Range module in XC Service Toolkit." Otherwise, select NA.)
16. Normalized Noise Power Spectrum (NNPS)
17. Detective Quantum Efficiency (DQE)
18. Image Plate Erasure (Erasure Thoroughness)
19. Repeated Exposure Test
20. Dynamic Range
21. Radiation Output
Radiation output rate is >800 mR/sec (7.0 mGy/sec) at 28 kVp with Mo/Mo
22. Sensitometric Response
23. Scanner Response Linearity
24. CR Cassette Uniformity
25. Diagnostic Review Workstation (RWS) QC (for all RWS, even if located offsite; NA if only hardcopy read)
26. DICOM Printer QC (Mammography Equipment Evaluations only)
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on M22: Click in boxes to use drop-down lists
Note on G29: Click in boxes to use drop-down lists

## Tech QC Eval-iCRco
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(iCRco, continued)
Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.
|  | FREQUENCY | PASS/FAIL
1. | Detector Warm-up | Daily
2. | Screen Cleaning/Erasure Test | Weekly
3. | SNR/CNR | Weekly
4. | Phantom Image Quality Evaluation | Weekly
5. | Visual Checklist | Weekly
6. | Repeat Analysis | Quarterly
7. | Compression | Semi-annually
8. | Image Plate Fog Test | Semi-annually
9. | Diagnostic Review Workstation QC (NA if only hardcopy read) | See FDA guidance
10. | Printer QC | See FDA guidance
Medical Physicist's Recommendations for Quality Improvement
Important: 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. 2. Use any QC manual version that is applicable to the model surveyed. 3. If the RWS or printer is FDA-cleared for FFDM, their QC manual is considered to be "substantially the same" and may be followed. (Check with the RWS or printer manufacturers for their clearance status and QC manual.) 4. If the RWS or printer is not cleared by the FDA for FFDM, follow the QC manual provided by the image receptor manufacturer. (Check with the image receptor manufacturer for their required tests.) 5. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day than the survey date, note the date above. 6. See the FDA-approved alternative standard for iCRco FFDM regarding corrective action periods when components fail QC. However, if these tests are performed as part of a Mammography Equipment Evaluation (e.g., for a new system), corrective action must be taken before mammographic images are acquired.
Note on J10: Click in boxes to use drop-down lists

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5