## TOC
Digital Mammography Quality Control Tests
Medical Physicist's Tests (2D and DBT)
Important: Before a facility may start using the procedures in the ACR Digital Mammography QC Manual for the first time on a unit, the medical physicist must first conduct an annual survey of the digital mammography unit and display devices using the manual and the ACR Digital Mammography Phantom.
Note: Complete Facility, Unit and Test Equipment Data tab first to populate facility information into forms
| Test | Minimum Frequency | Corrective Action Timeframe*
1. | Mammography Equipment Evaluation - MQSA Requirements | MEE | Before clinical use
2. | ACR DM Phantom Image Quality | MEE and Annual | Before clinical use
3. | DBT Z Resolution | MEE and Annual | Within 30 days
4. | Spatial Resolution | MEE and Annual | Within 30 days
5. | DBT Volume Coverage | MEE and Annual | Before clinical use
6. | Automatic Exposure Control System Performance | MEE and Annual | Within 30 days
7. | Average Glandular Dose | MEE and Annual | Before clinical use
8. | Unit Checklist | MEE and Annual | Critical: before clinical use; less critical: w/in 30 days
9. | Computed Radiography (if applicable) | MEE and Annual | Before clinical use
10. | Acquisition Workstation (AW) Monitor QC | MEE and Annual | W/in 30 days; before clinical use for severe defects
11. | Radiologist Workstation (RW) Monitor QC | MEE and Annual | W/in 30 days; before clinical use for severe defects
12. | Film Printer QC (if applicable) | MEE and Annual | Before clinical use
13. | Evaluation of Site's Technologist QC Program | Annual | Within 30 days
14. | Evaluation of Display Device Technologist QC Program | MEE and Annual | Within 30 days
15. | Manufacturer Calibrations (if applicable) | MEE and Annual | Before clinical use
16. | Collimation Assessment | DBT-MEE and Annual | Within 30 days
DM-MEE and Troubleshooting
| MEE or Troubleshooting - Beam Quality (Half-Value Layer) Assessment | MEE or Troubleshooting | MEE - before clinical use; troubleshooting - w/in 30 days
| MEE or Troubleshooting - kVp Accuracy and Reproducibility | MEE or Troubleshooting | MEE - before clinical use; troubleshooting - w/in 30 days
| Troubleshooting - Ghost Image Evaluation | Troubleshooting | Before clinical use
| Troubleshooting - Viewbox Luminance | Troubleshooting | NA
* Corrective action for MEEs must be performed before clinical use.
Summary Report Forms
Medical Physicist's ACR DM QC Test Summary
Mammography Technique Chart
Medical Physicist QC Letter for the Radiologist
Supplemental Forms
Facility, Unit and Test Equipment Data
QC Equipment List - Medical Physicist
| ACR Digital Mammography (DM) Phantom | Dosimeter
| (may use facility's phantom) | Lead sheet or equivalent
| 0.1 mm aluminum sheets | Photometer to measure luminance
| Hi-resolution bar pattern - 2 to 10 lp/mm | Thin metal ruler (for CR)
| 2, 4, 6 and 8 cm thick acrylic, BR-12 or BR-50 sheets | Coins, ready-pack film, electronic collimation test tools, or equivalent
kV meter

## Instr-Rev
Medical Physicist Test Form Spreadsheet
Instructions
• | IMPORTANT NOTE: This Excel Workbook contains some formulae to analyze the data and logic to report pass/fail results in order to help the medical physicist. It is intended to be used with both Windows and Mac operating systems. Please note, ANY alteration of the formulae or logic will result in errors. It is critical that you review the calculated results for accuracy and pass/fail designations for applicability on each sheet as well as the summary. It is the responsibility of the medical physicist to check all automated calculations and pass/fail results in the spreadsheet for correctness. Please notify the ACR if you find any errors.
• | Complete facility information in Fac Info tab first; information auto-populates in other pages.
• | Drop down lists are available within cells to complete the required information; if the list does not contain the necessary information; complete the cell yourself.
• | Your system may have multiple, clinically-used Image Modes that must be tested (e.g., 2D, 2D w/Add-on DBT Device, DBT). See next tab for the required tests.
If completing via paper, write in the Image Mode that is being tested on the top of all applicable forms. Make additional copies of the forms for each image mode tested.
If completing via computer, make a copy of the tab with each form for all applicable image modes tested. Select the Image Mode tested on each form using the drop-down list of Image Modes.
• | Enter data into white cells.
• | Light yellow cells are for pass/fail results of subtests within the main test; depending on the page, you may enter the pass/fail results using the dropdown list OR the results may populate via a formula.
• | Bright yellow cells are for pass/fail results of the overall test; depending on the page, you may enter the pass/fail results using the dropdown list OR the results may populate via a formula.
• | Light orange cells contain formulae or pre-entered values - DO NOT enter data in these cells; if you do, the formulae and results will be lost.
• | Do not complete gray cells.
Revisions
| 2016-09-24 | New
| 2017-01-13 | Misc typo and formatting corrections
| 2017-08-11 | Instructions-Revisions tab moved from end to beginning of the spreadsheet
Test #5 Average Glandular Dose: Moved equipment setup figure to right of page
Test #5 Average Glandular Dose: Corrected s-factor for W/Ag from 1.072 to 1.042
Test #5 Average Glandular Dose: Added s-factor for Rh/Ag
Misc typo and formatting corrections
| 2018-11-19 | Updated for DBT
| 2019-09-30 | Minor formula corrections; typo correction in TOC (Tech QC review freq); added DBT AEC tab

## Modes
Required Tests for Imaging Modes Used on 2D and DBT Systems

## Fac Info
Facility, Unit and Test Equipment Data
Medical Physicist's Tests
Facility Information
Facility Name
Address
Address
City, State, Zip
MAP ID# (00000)
MAP Unit# (00)
Lead Interpreting Radiologist
Quality Control Technologist
Survey Information
Room ID
Survey Date
MP Report Date
Date of Previous Survey
Unit Information
Digital Radiography (DR)
Digital Breast Tomosynthesis (DBT)
Computed Radiography (CR)
X-Ray Unit Manufacturer
X-Ray Unit Model
X-Ray Unit Gantry Serial #
X-Ray Unit Date of Manufacture
X-Ray Unit Date of Installation
SID (cm)
DC Offset
mA Large
mA Small
Magnification Stand Factor Used
Nominal Pixel Size (µm)
CR Reader Manufacturer
CR Reader Model
CR Reader Serial #
CR Date of Manufacture
CR Date of Installation
Test Equipment Info
ACR DM Phantom Manufacturer and S/N
Dosimeter Manufacturer/Model
Dosimeter Calibration Date
kVp Meter Manufacturer/Model
kVp Meter Calibration Date
Medical Physicist Info
Medical Physicist Name
Telephone Number
email
Signature

## 1.
1. Mammography Equipment Evaluation (MEE)
| Facility Name | 0 |  | MAP ID-Unit# (00000-00) | 0 | - | 0
| Mfr & Model |  |  | Room ID | 0
|  |  |  | Survey Date | 0
MQSA Requirements for Equipment [FDA Rule Sec. 900.12 (b)] - only applies to MEE
Feature |  | FDA Rule | Requirement |  |  |  | Meets? Yes/No/NA
Motion of tube-image receptor assembly |  | 3(i) | The assembly shall be capable of being fixed in any position where it is designed to operate. Once fixed in any such position, it shall not undergo unintended motion.
|  | 3(ii) | This mechanism shall not fail in the event of power interruption.
Image receptor sizes |  | 4(iii) | Systems used for magnification procedures shall be capable of operation with the grid removed from between the source and image receptor.
Light fields |  | 5 | For any mammography system with a light beam that passes through the X-ray beam-limiting device, the light shall provide an average illumination of not less than 160 lux (15 ft-candles) at 100 cm or the maximum source-image receptor distance (SID), whichever is less.
Magnification |  | 6(i) | Systems used to perform noninterventional problem-solving procedures shall have radiographic magnification capability available for use by the operator.
|  | 6(ii) | Systems used for magnification procedures shall provide, at a minimum, at least one magnification value within the range of 1.4 to 2.0.
Focal spot selection |  | 7(i) | When more than one focal spot is provided, the system shall indicate, prior to exposure, which focal spot is selected.
|  | 7(ii) | When more than one target material is provided, the system shall indicate, prior to exposure, the preselected target material.
|  | 7(iii) | When the target material and/or focal spot is selected by a system algorithm that is based on the exposure or on a test exposure, the system shall display, after the exposure, the target material and/or focal spot actually used during the exposure.
Application of compression |  | 8(i)(A) | Each system shall provide an initial power-driven compression activated by hands-free controls operable from both sides of the patient.
|  | 8(i)(B) | Each system shall provide fine adjustment compression controls operable from both sides of the patient.
Compression paddle |  | 8(ii)(A) | Systems shall be equipped with different sized compression paddles that match the sizes of all full-field image receptors provided for the system.
|  | 8(ii)(B) | Compression paddle shall be flat and parallel to the breast support table and shall not deflect from parallel by more than 1.0 cm at any point on the surface of the compression paddle when compression is applied.
|  | 8(ii)(C) | Equipment intended by the manufacturer’s design to not be flat and parallel to the breast support table during compression shall meet the manufacturer’s design specifications and maintenance requirements.
|  | 8(ii)(D) | Chest wall edge of the compression paddle shall be straight and parallel to the edge of the image receptor.
|  | 8(ii)(E) | Chest wall edge may be bent upward to allow for patient comfort but shall not appear on the image.
Technique factor selection and display |  | 9(i) | Manual selection of mAs or at least one of its component parts (mA and/or time) shall be available.
|  | 9(ii) | The technique factors (kVp and either mA and seconds or mAs) to be used during an exposure shall be indicated before the exposure begins, except when AEC is used, in which case the technique factors that are set prior to the exposure shall be indicated.
|  | 9(iii) | Following AEC mode use, the system shall indicate the actual kVp and mAs (or mA and time) used during the exposure.
Lighting** |  | 14 | The facility shall make special lights for film illumination, i.e., hot-lights, capable of producing light levels greater than that provided by the view box, available to the interpreting physicians.
Film masking devices** |  | 15 | Film masking devices that can limit the illuminated area to a region equal to or smaller than the exposed portion of the film are available to all interpreting physicians interpreting for the facility.
Beam quality assessment |  | * | Must meet the specifications of FDA's Performance Standards for Ionizing Radiation Emitting Products (Part 1020.30)
kVp accuracy & reproducibility |  | * | The mean kVp must not differ from the nominal by more than + 5% of the nominal kVp.
|  | * | The coefficient of variation must be ≤0.02.
Collimation assessment |  | * | If sum of left plus right edge deviations or anterior plus chest edge deviations exceeds 2% of SID, seek service adjustment.
|  | * | If X-ray field exceeds image receptor at any side by more than + 2% of SID or if X-ray field falls within image receptor on the chest wall side, seek service adjustment.
|  | * | If chest-wall edge of compression paddle is within the image receptor or projects beyond the chest-wall edge of the image receptor by more than 1% of SID, seek service correction.
Overall Pass/Fail
** NA is acceptable if 1) no hard copy interpretations are made, 2) no hard copy comparisons are made or 3) for new units at existing facilities if these were previously evaluated and have not changed

## 2.
2. ACR DM Phantom Image Quality
|  | Facility Name | 0 | [9 empty] | MAP ID-Unit# (00000-00) | 0 |  | - | 0
|  | Mfr & Model | [10 empty] | Room ID | 0
[5 empty] | ACR DM Phantom Mfr and S/N | 0 | [6 empty] | Survey Date | 0
Phantom Setup |  |  | Equipment: |  | ACR DM Phantom (required) | [6 empty] | Phantom Setup: |  |  |  | AEC mode:
|  |  | Follow procedure in the Technologist's ACR Technique & Procedure Summaries: | [12 empty] | Paddle size (IR size):
|  |  | • Use clinical technique for typical screening exam of 4.2 cm 50/50 breast | [12 empty] | Paddle type (reg or flex):
|  |  | • Largest IR & paddle, 5 daN or 12 lbs, score on AW | [12 empty] | View or selected image:
|  |  | • Adjust W/L to optimize test objects, zoom & pan entire image | [12 empty] | Compression force:
[7 empty] | WL |  | WW | [6 empty] | AEC cell position (if avail):
|  |  | • For DBT, scroll to best slice/slab to visualize test objects | [9 empty] | Target/filter (if app): |  |  | kVp (if app):
[9 empty] | Slice/slab # | [6 empty] | Density setting (if app):
[8 empty] | Phantom patient name: | [7 empty] | Mag factor (mag mode only): | 0
[8 empty] | Phantom patient ID: | [7 empty] | Image sent to which PACS?
Image Mode (2D, 2D w/Add-on DBT, DBT) | [7 empty] | 2D | [8 empty] | Mag Mode 2D
Resulting Techniques (if available) |  |  | Target/filter
Image receptor size
kVp
mAs
Unit-indicated AGD (mGy)
ACR DM Phantom Evaluation |  |  | Artifacts P/F
Fiber score
Speck group score
Mass score
|  |  | Phantom P/F |  |  |  |  | F |  | F | F |  | F |  | F |  | F
SNR & CNR Raw Image (does not apply to DBT) |  |  | DC offset (if applicable) |  |  |  |  | 0 |  | CNR from MEE (if avail; does not apply to MEEs) | CNR Lower Limit (85% of MEE) | CNR ≥ -15% of MEE (P/F)
Mean cavity signal
Mean background signal
Std dev of background
|  |  | Calculated SNR |  |  |  |  | #DIV/0!
|  |  | Calculated CNR |  |  |  |  | #DIV/0!
|  |  | SNR ≥ 40.0 (P/F) |  |  |  |  | #DIV/0!
|  |  | CNR ≥ 2.0 (P/F) |  |  |  |  | #DIV/0! |  |  | 0 | #DIV/0!
Distance Measurement |  |  | Parallel to A-C axis (mm)
|  |  | Meas = 70.0 ±14.0 mm (P/F) |  |  |  |  | F |  | F | F |  | F |  | F |  | F
Overall Pass/Fail
Initiated (or updated) technologist's ACR Technique and Procedure Summaries form
Analysis
| Full Point |  |  | Half Point
Fibers | ≥ 8 mm long |  |  | ≥5 & <8 mm
Specks | 4 - 6 specks |  |  | 2 - 3 specks
Masses | ≥ ¾ border |  |  | ≥ ½ & < ¾ border
Action Limits |  |  | Required: |  | ACR DM Phantom image must be free of clinically significant artifacts.
Fiber score must be ≥ 2.0; speck group score must be ≥ 3.0; mass score must be ≥ 2.0.
2D Only - MEE & Annual: SNR must be ≥ 40.0; CNR ≥ 2.0. Annual: CNR must be ≥ 85% of MEE.
Measured wax insert distance must be 70.0 ± 14.0 mm.
|  |  | Timeframe: |  | Failures of required items must be corrected before clinical use.
Distance

## 3.
3. DBT Z Resolution
|  | Facility Name | 0 | [6 empty] | MAP ID-Unit# (00000-00) | 0 |  |  | - | 0
|  | Mfr & Model | [7 empty] | Room ID | 0
[10 empty] | Survey Date | 0
Procedure |  |  | Equipment: | DBT image from ACR DM Phantom Image Quality test | [6 empty] | MEE Date:
MEE Baseline FWHM:
| Slice # |  | Relative Slice Location | Maximum Speck Signal | [10 empty] | Mean Background Signal
|  |  |  | Center Speck | 12:00 Speck | 2:00 Speck | 5:00 Speck | 7:00 Speck |  | 10:00 Speck |  | Ave Max Signal |  |  |  |  | Half Max Location
|  |  | -3 | [8 empty] | #DIV/0! |  |  |  |  | 0 | 0.5 |  |  | 0
|  |  | -2 | [8 empty] | #DIV/0! |  |  |  |  | 0 | Rising edge |  |  | 0
|  |  | -1 | [8 empty] | #DIV/0! |  |  |  |  | 0
|  |  | 0 | [8 empty] | #DIV/0! |  |  |  |  | 0
|  |  | 1 | [8 empty] | #DIV/0! |  |  |  |  | 0 | Falling edge
|  |  | 2 | [8 empty] | #DIV/0! |  |  |  |  | 0
|  |  | 3 | [8 empty] | #DIV/0!
[9 empty] | Z-Res Diff (Ave Max - Background Mean) |  |  |  | △ Z-Res Diff relative to DBT Slice 0 |  |  |  | For Chart
[17 empty] | Loc | Val | FWHM
[9 empty] | #DIV/0! | [7 empty] | -3 |  | 0 | 0.5
[9 empty] | #DIV/0! | [7 empty] | -2 |  | 0 | 0.5
[9 empty] | #DIV/0! | [7 empty] | -1
[9 empty] | #DIV/0! | [7 empty] | 0
[9 empty] | #DIV/0! | [7 empty] | 1
[9 empty] | #DIV/0! | [7 empty] | 2
[9 empty] | #DIV/0! | [7 empty] | 3
[9 empty] | Current FWHM |  |  |  |  | 0.00 |  | mm
[9 empty] | MEE Baseline FWHM |  |  |  |  | 0 |  | mm
[9 empty] | % Change (Current vs MEE) |  |  |  |  | #DIV/0!
Overall Pass/Fail
Action Limits |  |  | Required: | Annual Survey: FWHM must be within ±30% of the baseline (MEE) value. MEE: No action limit.
|  |  | Timeframe: | Failures must be corrected within 30 days.
Slice -2
Slice -1
Slice 0
Slice 1
Slice 2
ROI Placement:
Mean Background Signal - Next to Center Speck
ROI Placement:
Maximum Signal - Over
Each Speck

## 4.
4. Spatial Resolution
|  | Facility Name | 0 | [5 empty] | MAP ID-Unit# (00000-00) | 0 | - | 0
|  | Mfr & Model | [6 empty] | Room ID | 0
[9 empty] | Survey Date | 0
Procedure |  |  | Equipment: |  |  | ACR DM Phantom, line-pair test tool |  |  |  |  | Phantom Setup:
|  |  | Place ACR DM Phantom reversed on breast support (wax insert away from chest edge) | [7 empty] | Paddle size (IR size):
|  |  | Place bar pattern on top of phantom and under paddle at ~45o | [7 empty] | Paddle type (reg or flex):
Lightly compress paddle to touch bar pattern
Acquire "raw" images using manual mode closest to ACR DM Phantom technique
| Image Mode (2D, 2D w/Add-on DBT, DBT) | [10 empty] | Mag Mode 2D
| Setup Techniques |  |  | Mag factor |  |  |  | Contact
Target/filter
kVp
mAs
| Spatial Resolution Score |  |  | Line-pair score
Overall Pass/Fail
Action Limits |  |  | Required: |  | For 2D, spatial resolution must be ≥ 4.0 lp/mm for contact mode and 6.0 lp/mm for magnification mode.
For DBT, spatial resolution must be ≥ 2.0 lp/mm for contact mode.
|  |  | Timeframe: |  | Failures must be corrected within 30 days; for MEEs, before clinical use.
Chest Edge
~45o

## 5.
5. DBT Volume Coverage
|  | Facility Name | 0 |  |  |  |  | MAP ID-Unit# (00000-00) | 0 | - | 0
|  | Mfr & Model | [5 empty] | Room ID | 0
[8 empty] | Survey Date | 0
Procedure |  |  | Equipment: |  | ACR DM Phantom, 2 sheets of 0.1 mm Al |  |  |  |  | Phantom Setup:
|  |  | Place ACR DM Phantom on breast support in the usual position | [6 empty] | Paddle size (IR size):
|  |  | Place Al sheets on top and bottom of DM phantom, diagonally across chest wall | [6 empty] | Paddle type (reg or flex):
Acquire DBT image of phantom
View reconstructed image and verify that both Al sheets are in focus within the volume
Contact Mode
| Setup Techniques | Mag factor |  |  |  |  | Contact
Target/filter
kVp
mAs
| Results (Yes/No/NA) | Lower Al sheet in focus within the volume
Upper Al sheet in focus within the volume
Overall Pass/Fail
Action Limits |  |  | Required: | Both sheets must be focused in volume
|  |  | Timeframe: | Failures must be corrected before clinical use.

## 6.2D.
6. Automatic Exposure Control System Performance (2D)
Image Mode (2D, 2D w/Add-on DBT, DBT)
|  | Facility Name | 0 | [11 empty] | MAP ID-Unit# (00000-00) | 0 |  |  |  | - | 0
|  | Mfr & Model | [12 empty] | Room ID | 0
[15 empty] | Survey Date | 0
Procedure |  |  | Equipment: | 2, 4, 6, 8 cm of BR-12, BR-50 or acrylic | [8 empty] | Phantom Setup: |  |  |  |  | Paddle size (IR Size):
|  |  | Install small paddle (reg or flex) (Use large if small not available) | [14 empty] | Paddle type (reg or flex):
|  |  | Use regular or flex paddle used for most clinical imaging | [14 empty] | AEC cell position (if avail):
|  |  | Set thickness at actual thickness of phantom (2, 4, or 6 cm) | [14 empty] | Mag setting: | 0
|  |  | Acquire images using clinical techniques | [14 empty] | Mfr DC offset, if app: | 0
|  |  | SNR data must be obtained from raw image | [14 empty] | Other settings:
Magnification stand, if used clinically for 2D
AEC Thickness Tracking
Mode | Thick-ness (cm) |  | Setup Techniques |  |  | Resultant Techniques | [5 empty] | Signal and Noise Measurements
|  |  | AEC Mode | Density setting |  | Target/ Filter | kVp |  | mAs | Other |  | Mean Bkgd Signal |  | Std Dev of Bkgd |  | DC Offset (if app) |  |  |  | SNR
Contact | 2 | [14 empty] | 0 |  |  |  | #DIV/0!
Contact | 4 | [14 empty] | 0 |  |  |  | #DIV/0!
Contact | 6 | [14 empty] | 0 |  |  |  | #DIV/0!
Contact | 8 | [14 empty] | 0 |  |  |  | #DIV/0!
Mag* | 4 | [14 empty] | 0 |  |  |  | #DIV/0!
*2D only
Analysis
Mode | Thick-ness (cm) |  | SNR |  | MEE and Annual | [5 empty] | Annual
[5 empty] | Lowest Limit for SNR* |  |  | Pass/Fail |  |  | MEE SNR |  |  |  | Lower Limit |  | Upper Limit |  | SNR within ±15% of MEE (P/F)
Contact | 2 |  | #DIV/0! | [11 empty] | 0 |  | 0 |  | #DIV/0!
Contact | 4 |  | #DIV/0! |  | 40 |  |  | #DIV/0! | [6 empty] | 0 |  | 0 |  | #DIV/0!
Contact | 6 |  | #DIV/0! | [11 empty] | 0 |  | 0 |  | #DIV/0!
Contact | 8 |  | #DIV/0! | [11 empty] | 0 |  | 0 |  | #DIV/0!
Mag* | 4 |  | #DIV/0! | [11 empty] | 0 |  | 0 |  | #DIV/0!
*2D only | [10 empty] | Overall Pass/Fail
Action Limits |  |  | Required: | MEE and Annual: SNR must be ≥ 40.0 for 4.0 cm in contact mode.
Annual: SNR must be within ±15% of MEE over the clinically used phantom thickness and imaging modes.
|  |  | Timeframe: | Failures must be corrected within 30 days; for MEEs, before clinical use.

## 6.DBT.
6. Automatic Exposure Control System Performance (DBT)
Image Mode (2D, 2D w/Add-on DBT, DBT)
|  | Facility Name | 0 | [11 empty] | MAP ID-Unit# (00000-00) | 0 |  |  |  | - | 0
|  | Mfr & Model | [12 empty] | Room ID | 0
[15 empty] | Survey Date | 0
Procedure |  |  | Equipment: | 2, 4, 6, 8 cm of BR-12, BR-50 or acrylic | [8 empty] | Phantom Setup: |  |  |  |  | Paddle size (IR Size):
|  |  | Install small paddle (reg or flex) (Use large if small not available) | [14 empty] | Paddle type (reg or flex):
|  |  | Use regular or flex paddle used for most clinical imaging | [14 empty] | AEC cell position (if avail):
|  |  | Set thickness at actual thickness of phantom (2, 4, or 6 cm) | [14 empty] | Mag setting: | 0
|  |  | Acquire images using clinical techniques | [14 empty] | Mfr DC offset, if app: | 0
|  |  | SNR data must be obtained from raw image | [14 empty] | Other settings:
Magnification stand, if used clinically for 2D
AEC Thickness Tracking
Mode | Thick-ness (cm) |  | Setup Techniques |  |  | Resultant Techniques | [5 empty] | Signal and Noise Measurements
|  |  | AEC Mode | Density setting |  | Target/ Filter | kVp |  | mAs | Other |  | Mean Bkgd Signal |  | Std Dev of Bkgd |  | DC Offset (if app) |  |  |  | SNR
Contact | 2 | [14 empty] | 0 |  |  |  | #DIV/0!
Contact | 4 | [14 empty] | 0 |  |  |  | #DIV/0!
Contact | 6 | [14 empty] | 0 |  |  |  | #DIV/0!
Contact | 8 | [14 empty] | 0 |  |  |  | #DIV/0!
Mag* | 4 | [14 empty] | 0 |  |  |  | #DIV/0!
*2D only
Analysis
Mode | Thick-ness (cm) |  | SNR |  | MEE and Annual | [5 empty] | Annual
[5 empty] | Lowest Limit for SNR* |  |  | Pass/Fail |  |  | MEE SNR |  |  |  | Lower Limit |  | Upper Limit |  | SNR within ±15% of MEE (P/F)
Contact | 2 |  | #DIV/0! | [11 empty] | 0 |  | 0 |  | #DIV/0!
Contact | 4 |  | #DIV/0! | [11 empty] | 0 |  | 0 |  | #DIV/0!
Contact | 6 |  | #DIV/0! | [11 empty] | 0 |  | 0 |  | #DIV/0!
Contact | 8 |  | #DIV/0! | [11 empty] | 0 |  | 0 |  | #DIV/0!
Mag* | 4 |  | #DIV/0! | [11 empty] | 0 |  | 0 |  | #DIV/0!
*2D only | [10 empty] | Overall Pass/Fail
Action Limits |  |  | Required: | MEE and Annual: SNR must be ≥ 40.0 for 4.0 cm in contact mode.
Annual: SNR must be within ±15% of MEE over the clinically used phantom thickness and imaging modes.
|  |  | Timeframe: | Failures must be corrected within 30 days; for MEEs, before clinical use.

## 7.
7. Average Glandular Dose
| Facility Name | 0 |  |  |  |  | MAP ID-Unit# (00000-00) | 0 |  |  | - | 0
| Mfr & Model | [5 empty] | Room ID | 0
|  |  | ACR DM Phantom Mfr & S/N | 0 |  |  | Survey Date | 0
Procedure |  | Equipment: |  | Dosimeter |  |  |  | Dosimetry system: | 0
|  |  |  | ACR DM Phantom |  |  |  | Calibration date: | 0
|  | Use the technique from the ACR DM Phantom image page. | [5 empty] | Correction factor, if app:
|  | Measure mR/mAs or total exposure for dose calculation(s). | [5 empty] | SID (cm): | 0
Make exposure measurements at 4.2 cm
Imaging Mode (2D, 2D w/Add-on DBT, DBT)
Technique Factors Resulting From ACR DM Phantom Acquisition |  | ACR DM Phantom equivalent breast thickness (cm) |  |  |  | 4.2
|  | ACR DM Phantom material |  |  |  | Acrylic
AEC mode
Target/filter
kVp
mAs
Exposure Data (at skin surface) |  | Measured HVL (mm Al)
mAs setting for manual exposure measurement
Exposure #1 (mR)
Exposure #2 (mR)
Exposure #3 (mR)
|  | Average exposure (mR) |  |  |  | #DIV/0! | #DIV/0! |  |  | #DIV/0! |  | #DIV/0! | #DIV/0!
|  | Exposure/mAs at skin entrance (mR/mAs) |  |  |  | #DIV/0! | #DIV/0! |  |  | #DIV/0! |  | #DIV/0! | #DIV/0! | DM Phantom g-factor * c-factor * 8.76 mGy/R Calculator
|  | Total exposure (mR) |  |  |  | #DIV/0! | #DIV/0! |  |  | #DIV/0! |  | #DIV/0! | #DIV/0! | Breast | HVL | Quad | Linear |  |  | Entrance Exp | g * c
AGD Calculation D = Kgcs |  | Average entrance exposure - K (mR) |  |  |  | #DIV/0! | #DIV/0! |  |  | #DIV/0! |  | #DIV/0! | #DIV/0! | Thickness (cm) | (mm Al) | Factor | Factor | Constant | g * c | Conversion | * 8.76 mGy/R
|  | g-factor x c-factor x (8.76 mGy/R) |  |  |  | 0.47654 | [7 empty] | 4.2 | 0 | 0.1539 | 0.3775 | 0.0544 | 0.0544 | 8.76 | 0.47654
s-factor
|  | Computed AGD (mGy) |  |  |  | #DIV/0! | #DIV/0! |  |  | #DIV/0! |  | #DIV/0! | #DIV/0! | Extra for additional measurements DM Phantom g-factor * c-factor * 8.76 mGy/R Calculator
AGD Result |  | Pass/Fail |  |  |  | #DIV/0! | [7 empty] | Breast | HVL | Quad | Linear |  |  | Entrance Exp | g * c
Indicated vs. Calculated AGD (if avail) |  | Unit-indicated AGD from DM Phantom image (mGy) | [11 empty] | Thickness (cm) | (mm Al) | Factor | Factor | Constant | g * c | Conversion | * 8.76 mGy/R
|  | % Difference |  |  |  | #DIV/0! | #DIV/0! |  |  | #DIV/0! |  | #DIV/0! | #DIV/0! | 4.2 |  | 0.1539 | 0.3775 | 0.0544 | 0.0544 | 8.76 | 0.47654
|  | Indicated within ±25% of measured? |  |  |  | #DIV/0!
Action Limits |  | Required: |  |  | AGD for a single cranio-caudal view of the ACR DM Phantom in either 2D or DBT mode must not exceed 3.0 mGy. | [8 empty] | Acrylic and BR-12 s-factors
[14 empty] | Target-Filter | s-factor
|  | Recommended: |  |  | If available, unit-indicated AGD should be within ±25% of calculated AGD. | [8 empty] | Mo/Mo | 1
|  | Timeframe: |  |  | Doses > 3 mGy must be corrected before clinical use; failures of the unit-indicated AGD must be corrected within 30 days. | [8 empty] | Mo/Rh | 1.017
[14 empty] | Rh/Rh | 1.061
[14 empty] | Rh/Al | 1.044
[14 empty] | Rh/Ag | 1.087
[14 empty] | W/Rh | 1.042
[14 empty] | W/Al (0.5 mm) | 1.134
[14 empty] | W/Al (0.7 mm) | 1.082
[14 empty] | W/Ag | 1.042
Acrylic g-factor * c-factor * 8.76 mGy/R Table
[14 empty] | Breast Thickness (cm) | HVL (mm Al)
[15 empty] | 0.3 | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6
[14 empty] | 2 | 2.944 | 3.301 | 3.639 | 3.945 | 4.226 | 4.49 | 4.72
[14 empty] | 4 | 1.672 | 1.897 | 2.114 | 2.348 | 2.589 | 2.82 | 3.071
[14 empty] | 6 | 1.164 | 1.32 | 1.471 | 1.639 | 1.781 | 2.015 | 2.22
[14 empty] | 8 | 0.847 | 0.967 | 1.087 | 1.195 | 1.315 | 1.483 | 1.647
BR-12 g-factor * c-factor * 8.76 mGy/R for BR-12 Table
[14 empty] | Breast Thickness (cm) | HVL (mm Al)
[15 empty] | 0.3 | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6
[14 empty] | 2 | 3.4164 | 3.7931 | 4.1435 | 4.4588 | 4.7567 | 5.0195 | 5.1421
[14 empty] | 4 | 1.8133 | 2.0586 | 2.2864 | 2.5316 | 2.7857 | 3.031 | 3.2762
[14 empty] | 6 | 1.1826 | 1.349 | 1.5067 | 1.6819 | 1.8746 | 2.0674 | 2.2864
[14 empty] | 8 | 0.8585 | 0.9811 | 1.1038 | 1.2264 | 1.349 | 1.533 | 1.7082

## 8.
8. Unit Checklist
|  | Facility Name | 0 |  |  | MAP ID-Unit# (00000-00) | 0 |  | - | 0
|  | Mfr & Model |  |  |  | Room ID | 0
[6 empty] | Survey Date | 0
Procedure |  |  | Equipment: | None
Inspect the unit and evaluate the functionality according to the checklist below
Item | [8 empty] | Yes/No/NA
1. | Free-standing unit is mechanically stable.*
2. | All moving parts move smoothly, without obstructions to motion.
3. | All locks and detents work properly.*
4. | Image receptor holder assembly is free from vibrations.*
5. | Image receptor slides smoothly into holder assembly (if applicable).
6. | Image receptor is held securely by assembly in any orientation (if applicable).*
7. | Patient or operator is not exposed to sharp or rough edges, or other hazards.*
8. | Paddles are all intact with no cracks or sharp edges.*
9. | Mammography area is clean and free from significant dust and debris that may cause artifacts.
10. | Operator protected during exposure by adequate radiation shielding.*
11. | All indicators working properly.
12. | Autodecompression can be overridden to maintain compression (and status displayed).*
13. | Manual emergency compression release can be activated in the event of a power failure.*
14. | Is the audible exposure indicator at an appropriate volume level?
15. | DBT assembly moves as designed through its range of motion.*
16. | Operator technique charts are current and posted.
17. | Other:
18. | Other:
19. | Other:
20. | Other:
Overall Pass/Fail
Action Limits |  |  | Required: |  | All items, both critical (*) and noncritical, must pass.
|  |  | Timeframe: |  | Failures of critical items (*) must be corrected before clinical use; less critical items must be corrected within 30 days.

## 9.
9. Computed Radiography (if applicable)
| Facility Name | 0 | [9 empty] | MAP ID-Unit# (00000-00) | 0 |  | - | 0
| Mfr & Model | [10 empty] | CR Room
|  |  | CR Reader Mfr & Model | [8 empty] | Survey Date | 0
|  |  | CR Serial Number | 0 | [7 empty] | Medical Physicist | 0
|  |  | CR Date of Manufacture | 0 | [7 empty] | Signature
Inter-Plate Consistency & Artifact Evaluation
Procedure |  | Equipment: |  |  | ACR DM Phantom |  |  |  |  | AEC mode: |  |  |  |  | AEC detector position:
|  | ACR DM Phantom on breast support plate with associated paddle. | [12 empty] | Target/filter:
|  | Auto-Time, Set kV to ACR DM Phantom kV, cell position 2 if available. | [12 empty] | kVp:
Small Cassettes
Cassette ID |  | mAs Evaluation |  |  |  | SNR Evaluation (if available) | [7 empty] | Artifact Pass/Fail | Overall Pass/Fail
|  | mAs |  |  | P/F | Signal | Std Dev | SNR |  |  | P/F
1 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
2 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
3 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
4 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
5 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
6 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
7 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
8 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
[5 empty] | Allowable | [6 empty] | Allowable
Minimum mAs: |  | 0 |  |  | #DIV/0! |  | Minimum SNR: |  | 0 |  |  | #DIV/0!
Mean mAs: |  | #DIV/0! |  |  |  |  | Mean SNR: |  | #DIV/0!
Maximum mAs |  | 0 |  |  | #DIV/0! |  | Maximum SNR |  | 0 |  |  | #DIV/0!
Small Cassettes
|  | mAs Evaluation |  |  |  | SNR Evaluation (if available) | [7 empty] | Artifact Pass/Fail | Overall Pass/Fail
Cassette ID |  | mAs |  |  | P/F | Signal | Std Dev | SNR |  |  | P/F
1 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
2 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
3 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
4 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
5 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
6 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
7 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
8 |  |  |  |  | #DIV/0! | [5 empty] | #DIV/0! |  |  |  | #DIV/0!
[5 empty] | Allowable | [6 empty] | Allowable
Minimum mAs: |  | 0 |  |  | #DIV/0! |  | Minimum SNR: |  | 0 |  |  | #DIV/0!
Mean mAs: |  | #DIV/0! |  |  |  |  | Mean SNR: |  | #DIV/0!
Maximum mAs |  | 0 |  |  | #DIV/0! |  | Maximum SNR |  | 0 |  |  | #DIV/0!
Action Limits |  | Required: |  |  | mAs must be within ±10% of average mAs.
SNR must be within ±15% of average SNR.
Must be free of clinically significant artifacts.
|  | Timeframe: |  |  | Failures must be corrected before clinical use.

## 9.cont
9. Computed Radiography (cont)
| Facility Name | 0 |  |  |  | MAP ID-Unit# (00000-00) | 0 | - | 0
| Mfr & Model |  |  |  |  | Room ID | 0
[6 empty] | Survey Date | 0
CR Reader Scanner Performance
Procedure |  | Equipment: | 2 thin metal rulers (or equivalent).
Place rulers in the shape of a T on a cassette which is placed on top of the breast support.
Expose at extremely low manual technique (~25 kVp, 4 mAs).
Pass/Fail
Parallel to Chest Wall
Perpendicular to Chest Wall
Action Limits |  | Recommendation: |  | The edges of the "T" should appear smooth and sharp. If they are not, and appear jagged or nonsmooth, then this could indicate a problem with the CR reader performance.
Breast Support
Cassette
Ruler
Ruler

## 10.
10. Acquisition Workstation (AW) Monitor QC
| Facility Name | 0 | [6 empty] | MAP ID-Unit# (00000-00) | 0 |  | - | 0
| Mfr & Model | [7 empty] | Room ID | 0
|  |  | Medical Physicist | 0 |  |  |  |  | Survey Date | 0
Signature
Procedure |  | Equipment: |  |  | Luminance meter
Note: Some of these QC tests may or may not be possible to perform depending on the monitor QC capabilities
Test Pattern Image Quality: Use TG18-QC, SMPTE or other relevant pattern (if available)
Luminance Check: TG18 LN8-01 & LN8-18 test patterns, or others that provide measure of Lmin & Lmax (if available)
|  |  | Monitor manufacturer: |  |  |  | Model: |  |  |  | Significant findings indicated on figure below
Monitor serial number
Monitor date of manufacture
Monitor Condition |  |  | Significant findings P/F
Test Pattern Image Quality (if available) |  |  | Test pattern centered appropriately?
0%-5% contrast boxes visible?
95%-100% contrast boxes visible?
Alphanumerics sharp and legible?
3 "Quality Control" patches visible (TG18)?
Line-pair images distinct (center)?
Line-pair images distinct (corners)?
Grayscale ramps smooth (if avail)?
|  |  | Test pattern P/F |  |  |  |  | Pass
Luminance Check (if available) |  |  | Measured Luminance minimum (cd/m2) | [7 empty] | Luminance Uniformity
|  |  | Mfr recommendation for Lmin (if avail) | [7 empty] | Center
|  |  | Lmin meets mfr recommendation ±30%? |  |  |  |  | #DIV/0! |  |  | Upper Left
|  |  | Measured Luminance maximum (cd/m2) | [7 empty] | Upper Right
|  |  | Mfr recommendation for Lmax (if avail) | [7 empty] | Lower Left
|  |  | Lmax meets mfr recommendation ±10%? |  |  |  |  | #DIV/0! |  |  | Lower Right
|  |  | Luminance check P/F | [7 empty] | Max |  | 0
DICOM GSDF |  |  | W/in ±10% of targeted contrast response P/F (if avail) | [7 empty] | Min |  | 0
Mfr Automated Test |  |  | Most recent set of mfr automated tests P/F | [7 empty] | % Diff |  | #DIV/0!
Overall Pass/Fail | [10 empty] | P/F |  | #DIV/0!
Action Limits |  | Required: |  |  |  | Any identified screen blemish that could interfere with clinical information must be removed.
Test pattern image quality must pass all visual tests.
Lmin must be within ±30% of mfr specifications (or, if not available ≤ 1.5 cd/m2).
Lmax must be within ±10% of mfr specifications (or, if not available ≥ 150 cd/m2).
Luminance uniformity must be ≤ 30%
GSDF measured contrast response must be within ±10% of targeted contrast response.
Mfr's automated tests must pass mfr specifications (if 1 test fails, indicate "F").
|  | Timeframe: |  |  |  | Significant monitor cleanliness defects must be corrected before clinical use; all other required tests must be corrected within 30 days.

## 11.
11. Radiologist Workstation (RW) Monitor QC
| Facility Name | 0 | [7 empty] | MAP ID-Unit# (00000-00) | 0 |  |  | - | 0
| Workstation ID | [8 empty] | Survey Date | 0
|  |  | Medical Physicist | 0 | [5 empty] | Signature
Procedure |  |  | Equipment: |  |  | ACR DM Phantom Image, luminance meter
Note: Some of these QC tests may or may not be possible to perform depending on the monitor QC capabilities
ACR DM Phantom: use phantom acquired from any DM within facility network, preferably one MP has acquired
Test Pattern Image Quality: Use TG18-QC, SMPTE or other relevant pattern
Luminance: TG18 LN8-01, LN8-18 & TG18 UNL80 test patterns or other relevant test patterns
|  |  | Monitor manufacturer: |  |  |  | Model: | Left* | Right*
Monitor serial number
Monitor date of manufacture
Ambient Light |  |  |  | Are ambient light conditions adequate for DM? | [7 empty] | Significant findings indicated on figures below
Monitor Condition |  |  |  | Significant findings P/F
ACR DM Phantom Evaluation |  |  |  | Artifacts P/F
Fiber score
Speck group score
Mass score
|  |  |  | Phantom P/F |  |  |  | F | F
Distance Measurement |  |  |  | Parallel to A-C axis (mm)
|  |  |  | Meas = 70.0 ±14.0 mm (P/F) |  |  |  | F | F
Test Pattern Image Quality |  |  |  | Test pattern centered appropriately?
0%-5% contrast boxes visible?
|  |  |  | 95%-100% contrast boxes visible? | [7 empty] | *Left and right monitors; complete additional forms if more than 2 monitors used
Alphanumerics sharp and legible?
3 "Quality Control" patches visible (TG18)?
|  |  |  | Line-pair images distinct (center)? | [8 empty] | Luminance Uniformity
|  |  |  | Line-pair images distinct (corners)? | [8 empty] | Monitor | Left |  | Right
|  |  |  | Grayscale ramps smooth? | [8 empty] | Center
|  |  |  | Test pattern P/F |  |  |  | F | F |  |  |  | Upper L
Luminance Check |  |  |  | Measured Luminance minimum (cd/m2) | [8 empty] | Upper R
|  |  |  | Mfr recommendation for Lmin (if avail) | [8 empty] | Lower L
|  |  |  | Lmin meets mfr recommendation ±30%? |  |  |  | #DIV/0! | #DIV/0! |  |  |  | Lower R
|  |  |  | Measured Luminance maximum (cd/m2) | [8 empty] | Max | 0 |  | 0
|  |  |  | Mfr recommendation for Lmax (if avail) | [8 empty] | Min | 0 |  | 0
|  |  |  | Lmax meets mfr recommendation ±10%? |  |  |  | #DIV/0! | #DIV/0! |  |  |  | % Diff | #DIV/0! |  | #DIV/0!
|  |  |  | Luminance check P/F | [8 empty] | P/F | #DIV/0! |  | #DIV/0!
DICOM GSDF (if avail) |  |  |  | W/in ±10% of targeted contrast response P/F | [8 empty] | Luminance Matching
Mfr Automated Test |  |  |  | Most recent set of mfr automated tests P/F | [8 empty] | P/F | #DIV/0!
Overall Pass/Fail
Action Limits |  | Required: |  |  | Any identified monitor blemish that could interfere with clinical information must be removed.
ACR DM Phantom image must be free of clinically significant artifacts.
Fiber score must be ≥ 2.0; speck group score must be ≥ 3.0; mass score must be ≥ 2.0.
Measured distance of wax insert must be 70.0 ±14.0 mm.
Test pattern image quality must pass all visual tests.
Lmin must be within ±30% of mfr specifications (or, if not available ≤ 1.5 cd/m2).
Lmax must be within ±10% of mfr specifications (or, if not available ≥ 420 cd/m2).
Luminance uniformity must be ≤30%; luminance matching must be ≤ 20%.
GSDF measured contrast response must be within ±10% of targeted contrast response.
Mfr's automated tests must pass mfr specifications (if 1 test fails, indicate "F").
|  | Recommended: |  |  | Ambient light conditions should be appropriate for mammography; max of 45 lux is recommended.
|  | Timeframe: |  |  | Phantom must pass and significant monitor cleanliness defects must be corrected before clinical use; all other required tests must be corrected within 30 days.

## 12.
12. Film Printer QC (if applicable)
| Facility Name | 0 |  |  |  |  | MAP ID-Unit# (00000-00) | 0 |  | - | 0
| Printer ID | [5 empty] | Survey Date | 0
|  |  | Medical Physicist | 0 |  |  | Signature
Procedure |  | Applicability: If film printer is used clinically for mammography (i.e., for interpretation and to provide images to referring physicians and patients)
Equipment: Densitometer
Print an ACR DM Phantom image acquired from any DM unit within facility network, preferably one MP has just acquired.
Do not change window/level settings from acquired image prior to printing.
Print the phantom image from the workstation/computer typically used to print clinical films.
Dmax should be measured either at extreme left or right edge of film or at extreme non-chest wall edge.
[5 empty] | Film Printer Manufacturer |  |  |  | Film Printer Serial Number
[5 empty] | Film Printer Model |  |  |  | Film Printer Date of Manufacture
Workstation for printing | [8 empty] | DM ID or workstation ID
Film size
ACR DM Phantom |  |  | Artifacts P/F
Fiber score
Speck group score
Mass score
|  |  | Phantom P/F |  |  | F |  | F
Back-ground |  |  | Bkgd OD (Outside cavity)
|  |  | Bkgd OD ≥ 1.6 (P/F) |  |  | F |  | F
Contrast |  |  | Cavity OD
Bkgd OD (use value from above)
|  |  | Contrast = Cavity OD - Bkgd OD |  |  | #VALUE! |  | #VALUE!
|  |  | Contrast ≥ 0.1 (P/F) |  |  | #VALUE! |  | #VALUE!
Dmax |  |  | Dmax OD
|  |  | Dmax OD ≥ 3.1 (P/F) |  |  | F |  | F
Distance Measurement |  |  | Parallel to A-C axis (mm)
|  |  | Meas = 70.0 ±14.0 mm (P/F) |  |  | F |  | F
Overall Pass/Fail | [5 empty] | #VALUE! |  | #VALUE!
Action Limits |  | Required: |  |  | ACR DM Phantom image must be free of clinically significant artifacts.
Fiber score must be ≥ 2.0; speck group score must be ≥3.0; mass score must be ≥ 2.0.
Background OD must be ≥ 1.6 (1.7 to 2.2 is recommended; approx 2.0 is optimal).
Contrast (Cavity OD - Background OD) must be ≥ 0.1.
Dmax must be ≥ 3.1 (≥ 3.5 is recommended).
Measured distance of wax insert must be 70.0 ±14.0 mm.
|  | Timeframe: |  |  | Failures of required items must be corrected before printing of clinical images.

## 13.
13. Evaluation of Site's Technologist QC Program
|  | Facility Name | 0 | [8 empty] | MAP ID-Unit# (00000-00) | 0 |  | - | 0
|  | Mfr & Model | [9 empty] | Room ID | 0
[12 empty] | Survey Date | 0
Radiologic Technologist's Quality Control Tests | [5 empty] | Frequency |  | Test Performed, Analyzed & Documented | Missing Data | Incorrect Scoring or Calculations | Missing Corrective Action Documentation | Other |  | Comments
1. | ACR DM Phantom Image Quality |  |  |  |  | Weekly
[9 empty] | Tech Score |  | MP Score
| Medical physicist comparison scores of latest phantom image: | [5 empty] | Fiber
Speck group
Mass
Artifacts
2. | CR Cassette Erasure (if app) |  |  |  |  | Weekly
3. | Comp Thickness Indicator |  |  |  |  | Monthly
4. | Visual Checklist |  |  |  |  | Monthly
5. | AW Monitor QC |  |  |  |  | Monthly
9. | Facility QC Review |  |  |  |  | Quarterly
10. | Compression Force |  |  |  |  | Semiannual
11. | Mfr Calibrations (if app)
Optional - Repeat Analysis | [5 empty] | As Needed
Optional - System QC for Radiologist | [5 empty] | NA
Optional - Radiologist IQ Feedback | [5 empty] | NA
Corrective Action Log documentation adequate?
Overall Pass/Fail for Performance of Technologist QC Program
Additional Comments:
Action Limits |  |  | Required: |  | MQSA regulations [FDA Rule 900.12(d)(1)(iii) specify that "each facility shall have the services of a medical physicist available to survey mammography equipment and oversee the equipment-related quality assurance practices of the facility. " Completion of this "Evaluation of Site's Technologist QC Program" form documents that this oversight has been conducted. In order for the overall evaluation to pass, there must be a) no significant missing data, b) the tests must be analyzed without gross errors, and c) appropriate corrective action for failures must be taken (and documented). See test procedures for more information.
|  |  | Timeframe: |  | Failures must be corrected within 30 days.

## 14.
14. Evaluation of Display Device Technologist QC Program
|  | Facility Name | 0 | [7 empty] | MAP ID-Unit# (00000-00) | 0 |  | - | 0
|  | Medical Physicist | 0 | [7 empty] | Display Device Location
|  | Signature | [8 empty] | Survey Date | 0
Display Device ID & Room |  |  |  |  | Display Device Description (RW, Printer, Viewbox) | Test Performed, Analyzed & Documented Incorrectly | Missing Data | Incorrect Scoring or Calculations | Missing Corrective Action Documentation | Mfr Automated Tests (if Applicable) | Other |  | Comments |  |  | P/F
Example: Mammography reading room |  |  |  |  | RW |  |  |  |  |  |  |  | Discussed with manager |  |  | P
Corrective Action Log documentation adequate?
Overall Pass/Fail for Performance of Display Device Technologist QC Program
Additional Comments:
Action Limits | Required: |  |  | MQSA regulations [FDA Rule 900.12(d)(1)(iii) specify that "each facility shall have the services of a medical physicist available to survey mammography equipment and oversee the equipment-related quality assurance practices of the facility. " Completion of this "Evaluation of Site's Technologist QC Program" form documents that this oversight has been conducted. In order for the overall evaluation to pass, there must be a) no significant missing data, b) the tests must be analyzed without gross errors, and c) appropriate corrective action for failures must be taken (and documented). See test procedures for more information.
| Timeframe: |  |  | Failures must be corrected within 30 days.

## 15.
15. Manufacturer Calibrations (if applicable)
|  | Facility Name | 0 |  |  |  | MAP ID-Unit# (00000-00) | 0 | - | 0
|  | Mfr & Model |  |  |  |  | Room ID | 0
[7 empty] | Survey Date | 0
Procedure |  |  | Follow manufacturer's instructions.
Notes:
| Name of Calibration |  |  |  |  | Detector
| Results (P/F) | 2D
2D w/Add-on DBT Device
DBT
Overall Pass/Fail
Action Limits |  |  | Required: | Unit must pass all manufacturer's calibrations to pass overall.
|  |  | Timeframe: | Failures must be corrected before clinical use.

## 16.
16. Collimation Assessment
Image Mode (2D, 2D w/Add-on DBT, DBT)
| Facility Name | 0 | [5 empty] | MAP ID-Unit# (00000-00) | 0 |  | - | 0
| Mfr & Model | [6 empty] | Room ID | 0
[8 empty] | Survey Date | 0
Procedure |  | Equipment: |  | Coins, film, electronic collimation test tool(s), etc.
|  |  | Eqpt used: | [5 empty] | kVp: |  | mAs:
[7 empty] | Largest Detector Size Available |  |  | Small Detector Size (CR only)
Technique |  |  |  |  | Target material
Collimator size (cm)
[5 empty] | SID (mm) |  | 0 |  |  | 0
Deviation Between X-ray Field and Light Field |  |  |  |  | Left edge deviation (mm)
Right edge deviation (mm)
[6 empty] | Sum of left and right edge deviations | 0 |  |  | 0
[5 empty] | Sum as % of SID |  | #DIV/0! |  |  | #DIV/0!
Anterior edge deviation (mm)
Chest edge deviation (mm)
[5 empty] | Sum of anterior and chest edge deviations |  | 0 |  |  | 0
[5 empty] | Sum as % of SID |  | #DIV/0! |  |  | #DIV/0!
Pass/Fail | [6 empty] | #DIV/0! |  |  | #DIV/0!
Deviation Between X-ray Field and Edges of the Image Receptor |  |  |  |  | Left edge deviation
[5 empty] | % of SID (retain sign) |  | #DIV/0! |  |  | #DIV/0!
Right edge deviation
[5 empty] | % of SID (retain sign) |  | #DIV/0! |  |  | #DIV/0!
Anterior edge deviation
[5 empty] | % of SID (retain sign) |  | #DIV/0! |  |  | #DIV/0!
Chest edge deviation
[5 empty] | % of SID (retain sign) |  | #DIV/0! |  |  | #DIV/0!
Pass/Fail | [6 empty] | #DIV/0! |  |  | #DIV/0!
Alignment of Chest-Wall Edges of Compression Paddle and IR |  |  |  |  | Difference between paddle edge and film
[5 empty] | Difference as % of SID |  | #DIV/0! |  |  | #DIV/0!
Pass/Fail | [6 empty] | #DIV/0! |  |  | #DIV/0!
Overall Pass/Fail
Action Limits |  | Required: |  | If sum of left plus right edge deviations or anterior plus chest edge deviations exceeds 2% of SID, seek service adjustment. If X-ray field exceeds image receptor at any side by more than +2% of SID or if X-ray field falls within image receptor on the chest wall side, seek service adjustment. If chest-wall edge of compression paddle is within the image receptor or projects beyond the chest-wall edge of the image receptor by more than 1% of SID, seek service correction.
|  | Timeframe: |  | Failures must be corrected within 30 days; for MEEs, before clinical use.

## MEE-TS HVL
MEE or Troubleshooting
Beam Quality (Half-Value Layer) Assessment
| Facility Name | 0 | [12 empty] | MAP ID-Unit# (00000-00) | 0 |  |  | - | 0
| Mfr & Model | [13 empty] | Room ID | 0
[15 empty] | Survey Date | 0
Procedure |  | Equipment: |  |  | Dosimeter, 0.1 mm Al sheets, lead sheet | [10 empty] | Dosimetry system: | 0
|  | Cover the detector with lead sheet or apron | [13 empty] | Calibration date: | 0
Make at least 1 measurement for each available target-filter combination used for 2D and DBT (as applicable)
[6 empty] | Target/Filter 1 |  | Target/Filter 2 |  |  | Target/Filter 3 |  |  | Target/Filter 4 |  |  | Target/Filter 5 |  |  | Target/Filter 6
Target/filter
Nominal kVp setting
mAs
Exposure Measurements
[6 empty] | mm AL | X (mR) | mm AL |  | X (mR) | mm AL | X (mR) |  | mm AL | X (mR) |  | mm AL | X (mR) |  | mm AL | X (mR)
No Aluminum |  |  |  | E0 |  | 0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
Al Thickness (mm) ta |  |  |  | Ea
Al Thickness (mm) tb |  |  |  | Eb
Calculated or measured HVL (mm Al) | [5 empty] | #DIV/0! |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0!
Minimum allowed HVL | [5 empty] | 0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
Overall Pass/Fail | [5 empty] | #DIV/0! |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0!
[8 empty] | HVL = | tb ln[2Ea/E0] - ta ln[2Eb/E0]
ln [Ea/Eb]
Action Limits |  | Required: |  |  | The HVL must meet the specifications of FDA's Performance Standards for Ionizing Radiation Emitting Products (Part 1020.30) as shown below.
|  | Timeframe: |  |  | All failures must be corrected before clinical use.
FDA X-ray Tube Voltage (kilovolt peak) and Minimum HVL
|  |  | Designed Operating Range (kV) | [5 empty] | Measured Operating Voltage (kV) |  |  |  | Minimum HVL (mm of Al)
|  |  | Below 50 | [5 empty] | 20 |  |  |  | 0.2
[9 empty] | 25 |  |  |  | 0.25
[9 empty] | 30 |  |  |  | 0.3

## MEE-TS kVp
MEE or Troubleshooting
kVp Accuracy and Reproducibility
|  | Facility Name | 0 | [6 empty] | MAP ID-Unit# (00000-00) | 0 |  | - | 0
|  | Mfr & Model | [7 empty] | Room ID | 0
[10 empty] | Survey Date | 0
kVp Accuracy and Reproducibility
Procedure |  |  | Equipment: | kVp meter, lead sheet
Cover the entire detector with lead sheet, a lead apron or other device.
Remove the paddle.
|  |  |  | kVp meter | 0 |  |  | Setting
|  |  |  | Calibration Date: | 0
[6 empty] | Low Clinical kVp* | ACR DM Phantom Clinical kVp (2D) |  | ACR DM Phantom Clinical kVp (DBT) |  |  | High Clinical kVp*
Technique | Nominal kVp setting
Target/filter
Focal spot
mAs
Data | Measured kVp value 1
Measured kVp value 2
Measured kVp value 3
Analysis | Mean kVp |  |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! |  |  | #DIV/0!
| Standard deviation (SD) | [5 empty] | #DIV/0!
| Mean kVp - nominal kVp |  |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! |  |  | #DIV/0!
| 0.05 x nominal kVp |  |  |  |  | 0 | 0 |  | 0 |  |  | 0
| % error |  |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! |  |  | #DIV/0!
| % error P/F |  |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! |  |  | #DIV/0!
| Coefficient of variation (CV) | [5 empty] | #DIV/0!
| CV P/F | [5 empty] | #DIV/0!
Overall Pass/Fail | [5 empty] | #DIV/0! | #DIV/0! |  | #DIV/0! |  |  | #DIV/0!
* The low and high kVps may be either those used for 2D or DBT imaging, whichever are lowest or highest
Action Limits |  |  | Required: | Mean kVp must not differ from the nominal by more than ±5% of the nominal kVp.
Coefficient of variation must be ≤ 0.02.
|  |  | Timeframe: | Failures must be corrected within 30 days; for MEEs, before clinical use.

## TS Ghost
Troubleshooting
Ghost Image Evaluation
[8 empty] | Image Mode | 2D
| Facility Name | 0 | [5 empty] | MAP ID-Unit# (00000-00) | 0 |  |  |  | - | 0
| Mfr & Model | [6 empty] | Room ID | 0
[8 empty] | Survey Date | 0
Procedure |  | Equipment: |  | ACR DM Phantom, 0.1 mm Al sheet (10 cm x 10 cm) | [6 empty] | Phantom Setup
|  | Largest image receptor size | [8 empty] | Paddle Size (IR size):
|  | Clinical paddle (reg or flex) | [8 empty] | Paddle Type (reg or flex):
|  | Apply 5 daN or 12 lbs comp force | [8 empty] | Exposure Mode:
[11 empty] | Compression Force: | 12 lbs or 5 daN
|  | Position ACR DM Phantom with wax insert opposite from chest wall edge. | [8 empty] | AEC Cell Position (if avail):
|  | AEC cell position to position "3". | [8 empty] | Density Setting:
Use clinical (AEC) technique for both images
Image 1: Position phantom like Setup Image #1 (edge extends 1" beyond midline).
Image 2: Position phantom like Setup Image #2 with Al placed on top.
Signal data must be obtained from raw image
[7 empty] | Image 1 |  |  | Image 2
Resulting Techniques from Image Acquisition |  |  |  | Target/filter
kVp
mAs
Ghosting Analysis (see images below) |  |  |  | S1
S2
S3
|  |  |  | Ghosting Index | [5 empty] | #DIV/0!
[6 empty] | Overall Pass/Fail |  |  |  | #DIV/0!
|  |  |  | Ghosting Index = | (S3 - S2)
(S1 - S2)
Action Limits |  | Required: | The ghosting index must be within 0±0.3.
|  | Timeframe: | Failures must be corrected before clinical use.
Setup Image 1
Setup Image 2
Evaluation Image
Image 1
Image 2
S1
S2
S3

## TS View
Troubleshooting
Viewbox Luminance
|  |  | Facility Name | 0 | [8 empty] | MAP ID-Unit# (00000-00) | 0 |  | - | 0
[5 empty] | Medical Physicist | 0 | [6 empty] | Viewbox Location
[5 empty] | Signature | [7 empty] | Survey Date | 0
Procedure |  |  |  |  | Equipment: |  | Luminance meter
Measure luminance for all viewboxes, record the luminance value for the viewbox with the lowest luminance.
Note: Only check a deficiency if it is significant and could impact interpretation; if the observation is not significant, just make a note in comments
Viewbox Designation | [8 empty] | Measurements | Significant Deficiencies | [5 empty] | Pass/Fail
[9 empty] | Viewbox Luminance (cd/m2) | Dirt and Marks | Color Difference | Luminance Difference |  | Non-Uniformity | Functioning Masks Missing
1.
2.
3.
4.
5.
6.
7.
8.
9.
10.
Tech QC Review for Viewbox Luminance
Overall Pass/Fail
Comments:
| Action Limits |  |  |  | Recommended: |  |  | Mammography viewboxes should be capable of a luminance of 3,000 cd/m2, be uniform, clean and have functioning masks; if these are not met, corrective action should be taken.

## MP-Sum
Medical Physicist's ACR DM QC Test Summary
Facility Name | 0 | [9 empty] | MAP ID-Unit# | 0 |  | - | 0
Address | 0 | [11 empty] | Room ID | 0
| 0 | [11 empty] | Report Date | 0
| 0 | [11 empty] | Survey Date | 0
|  | X-Ray Unit Manufacturer | 0 | [6 empty] | Model | 0
|  | Control Panel Serial # | 0 | [5 empty] | Manufacture Date | 0 |  |  |  | Installation Date | 0
DM Unit Type: |  |  | Digital radiography (DR) |  | Computed radiography (CR) | [5 empty] | Digital Breast Tomosynthesis (DBT)
| Unit Use: |  | Diagnostic and screening mammography |  |  |  |  | Diagnostic only |  |  |  | Screening only
Survey Type: |  |  | Mammography equipment evaluation (MEE) - Full | [5 empty] | MEE - Partial |  |  |  |  | Annual survey
Equipment Tested: |  |  | DM unit | AW monitor |  | RW monitor |  |  | Viewbox |  |  | Printer | Other:
Oversight Level: |  |  | Medical physicist on-site |  |  | Medical physicist oversight
Quality Control Manual Used for Survey and Facility QC: | [5 empty] | 2018 ACR Digital Mammography QC Manual (with 2D and DBT QC)
Medical Physicist |  | 0 |  |  |  | Signature
QC Test Results
Test |  |  |  |  | Pass/Fail* | [6 empty] | CA
[5 empty] | 2D** |  | 2D Add-on DBT |  |  | DBT
Medical Physicist Tests
1. Mammography Equipment Evaluation - MQSA Reqs
2. ACR DM Phantom Image Quality | [12 empty] | Your Phantom Results - 2D
3. DBT Z Resolution
4. Spatial Resolution | [14 empty] | Fiber (≥ 2.0)
5. DBT Volume Coverage | [14 empty] | Speck grp (≥ 3.0)
6. Automatic Exposure Control System Performance | [14 empty] | Mass (≥ 2.0)
7. Average Glandular Dose | [14 empty] | AGD (≤ 3.0 mGy)
8. Unit Checklist
9. Computed Radiography (if applicable) | [12 empty] | Your Phantom Results - DBT
10. Acquisition Workstation Monitor QC
11. Radiologist Workstation Monitor QC | [14 empty] | Fiber (≥ 2.0)
12. Film Printer QC (if applicable) | [14 empty] | Speck grp (≥ 3.0)
13. Evaluation of Site's Technologist QC Program | [14 empty] | Mass (≥ 2.0)
14. Evaluation of Display Device Technologist QC Program | [14 empty] | AGD (≤ 3.0 mGy)
15. Manufacturer Calibrations (if applicable)
16. Collimation Assessment
MEE/Troubleshooting - Beam Quality (HVL) Assessment
MEE/Troubleshooting - kVp Accuracy and Reproducibility
Troubleshooting - Ghost Image Evaluation
Troubleshooting - Viewbox Luminance
Technologist QC Evaluation | [8 empty] | Date reviewed if after new unit MEE:
1. ACR DM Phantom Image Quality
2. Computed Radiography Cassette Erasure (if applicable)
3. Compression Thickness Indicator
4. Visual Checklist
5. Acquisition Workstation Monitor QC | [12 empty] | * "Pass" means all components of test passes; "Fail" means any or all components fail; if "CA" checked, see Corrective Action Summary ** or DBT aquistition only
6. Radiologist Workstation Monitor QC
7. Film Printer QC (if applicable)
8. Viewbox Cleanliness (if applicable)
9. Facility QC Review
10. Compression Force
11. Manufacturer Calibration (if applicable)
Optional - Repeat Analysis

## MP-Sum cont
Medical Physicist's ACR DM QC Test Summary (cont)
|  | Facility Name | 0 |  | MAP ID-Unit# (00000-00) | 0 |  | - | 0
|  | Mfr & Model |  |  | Room ID | 0
[5 empty] | Survey Date | 0
Corrective Action Summary*
*Note: This is only a summary page, the Corrective Action Log Form may contain further details.
Utilize Corrective Action Log Form
Required/ Recommended |  |  | Time Frame | Description |  |  | Date Completed |  |  | Initials

## T-Chart
Mammography Technique Chart
Image Mode (2D, 2D w/Add-on DBT, DBT)
|  | Facility Name | 0 |  |  | MAP ID-Unit# (00000-00) | 0 |  | - | 0
|  | Mfr & Model |  |  |  | Room ID | 0
[6 empty] | Survey Date | 0
Screening/Diagnostic Digital Mammography
| Compressed Breast Thickness |  |  | 50% Fatty - 50% Dense Breast
|  |  |  | AEC Mode | Target/Filter |  | kVp
< 3 cm
3 to 5 cm
5 to 7 cm
> 7 cm
.
Implant Displaced Mammography Views (Manual Technique)
| Breast Size |  |  | Target/Filter | kVp |  | mAs
Small
Medium
Large
ACR DM Phantom Technique (Weekly QC)
|  |  |  | Digital Mammography | DBT |  | 2D w/DBT Fixture
AEC mode
Paddle size (IR Size)
Paddle type (reg or flex)
View/selected image type
Slice or Slab # (DBT only)
Compression force
AEC cell position (if avail)
Target/filter (if app)
kVp (if app)
Density setting (if app)

## Rad Ltr
Medical Physicist QC Letter for the Radiologist
0
0
0
0
0
0
Re: Medical Physicist Survey of | [5 empty] | on | 0
Dear Lead Interpreting Radiologist,
The above mammography unit at your facility recently underwent an Annual Medical Physics Survey. Below is the relevant summary information as a result of this survey. Please note that your facility must follow-up on the Action Items below and obtain relevant documentation from the service engineer. Please evaluate the ACR Digital Mammography Phantom image acquired during the medical physicist testing (Image ID information listed below) and see my comments. If you have any questions please don't hesitate to call.
● Image Quality
|  | Patient Name (Phantom): | 0
|  | Patient ID (Phantom): | 0
|  | Date: | 0
ACR Digital Mammography Phantom Scores
0
|  |  | 2D | DBT | Passing Criteria | Pass /Fail
|  | Fiber score |  |  | ≥ 2.0
|  | Speck group score |  |  | ≥ 3.0
|  | Mass score |  |  | ≥ 2.0
|  | Artifacts |  |  | No Clinically Significant Artifacts
Comments on phantom image:
● Radiation Dose
ACR Digital Mammography Phantom Radiation Dose Values
0
|  |  | 2D | DBT | Passing Criteria | Pass /Fail
|  | ACR Phantom Dose (mGy) |  |  | ≤ 3.0
|  | Note: | The above dose is an estimate determined with a phantom representing the FDA-defined 4.2 cm thick, 50% glandular/50% adipose standard breast. Doses will vary with patient size and density. Specific patient doses can be estimated by your medical physicist.
Comments on radiation dose:

## Rad Ltr cont
Medical Physicist QC Summary Letter for the Radiologist (cont)
● Required Action Items
|  | Time Frame | Description
● Recommended Action Items
|  | Time Frame | Description
● Comments on Monitors, Monitor QC, & Viewing Conditions
|  | Time Frame | Description
● Comments on Tech QC
|  | Time Frame | Description
If you have any questions, please do not hesitate to call.
Sincerely,
|  |  |  | Phone | 0
0 |  |  |  | Email | 0

## DD
PF/PFNA | Pass/Fail
P | Pass
F | Fail
NA | NA
TF | Paddle
Mo/Mo | reg
Mo/Rh | flex
Rh/Al | other
Rh/Ag
Rh/Rh
W/Ag | None
W/Al | None
W/Rh
Target
Molybdenum
Rhodium
Tungsten
YN
Yes
No
NAYN
NA
Yes
No
Acrylic?
acrylic
BR-12
Check

Y/N/NA
Yes
No
NA
Timeframe
Immediately
W/in 30 days
NA
Req
Required
Recommended
NA
Display Device
RW
Printer
Viewbox
DM-DBT
2D
2D w/Add-on DBT
DBT
All