## QC Test Summary-GE Pristina
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – General Electric Pristina
Site Name | [7 empty] | Report Date
Address | [7 empty] | Survey Date
Medical Physicist's Name | [7 empty] | Signature
X-Ray Unit Manufacturer |  |  | General Electric |  |  |  |  | Model | Pristina
Date of Installation | [7 empty] | Room ID
QC Manual Version # | [5 empty] | (use version applicable to unit tested; contact mfr if questions)
Accessory Equipment |  |  | Manufacturer |  | Model |  | Location |  | QC Manual Version #
Review Workstation*
Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System (www.accessdata.fda.gov/cdrh_docs/presentations/pghs/Polic_Guidance_Help_System.htm).
Survey Type |  | Mammo Eqpt Evaluation (MEE) of new unit (include MQSA Rqmts for Mammo Eqpt checklist) | [8 empty] | Annual Survey
Features |  | 2D |  | Digital Breast Tomosynthesis (DBT)
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
PASS/FAIL
1. Acquisition Station Monitor Cleaning and Check
2. Image Uniformity and Bad Pixels (Flat Field) Test (NA GE Pristina 5845677-8EN, Rev. 2)
3. Artifact Evaluation and Flat Field Uniformity
4. IQST Test
5. Phantom Image Quality |  |  |  |  | Fibers | Specks | Masses
Phantom IQ Test on AWS
Phantom IQ Test on Printer (if applicable)
6. AOP and SNR Check
7. Collimation Assessment (MEE and troubleshooting only GE Pristina 5845677-8EN, Rev. 2)
8. Compression Paddles Chest Wall Edge Alignment (MEE or troubleshooting)
9. Sub-System MTF Measurement (NA GE Pristina 5845677-8EN, Rev. 2)
10. Breast Entrance Exposure, Average Glandular Dose and Reproducibility
| Average glandular dose for average breast is ≤3 mGy (300 mrad) | [8 empty] | mrad
Exposure reproducibility (CV) for air kerma (R) and mAs is ≤0.05
11. Test for Flexible Paddle Deflection in Compression (MEE, Annually and troubleshooting, GE Pristina 5845677-8EN, Rev. 2)
12. kVp Accuracy and Reproducibility (MEE and troubleshooting only)
13. Beam Quality Assessment (Half-Value Layer Measurement) (MEE and troubleshooting only GE Pristina 5845677-8EN,Rev. 2)
14. Radiation Output (NA GE Pristina 5845677-8EN, Rev. 2)
15. Mammographic Unit Assembly Evaluation
Meets requirements for motion of tube-image receptor assembly
Meets requirements for compression paddle decompression
16. Reconstructed Image Uniformity Test (DBT) (NA GE Pristina 5845677-8EN, Rev. 2)
17. Artifact Evaluation and Flat Field Uniformity (DBT) (GE Pristina 5845677-8EN, Rev. 2)
18. Phantom Image Quality - 3D (DBT) |  |  |  |  | Fibers | Specks | Masses
Phantom IQ 3D Test
19. AOP 3D Check (DBT) (NA GE Pristina 5845677-8EN, Rev. 2)
20. Breast Entrance Exposure and Average Glandular Dose - 3D (DBT)
| Average glandular dose for average breast is ≤3 mGy (300 mrad) | [8 empty] | mrad
21. Volume Coverage (DBT)
22. Review Workstation (RWS) Tests (for all RWS, even if located offsite; NA if only hardcopy read)
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on N22: Click in boxes to use drop-down lists
Note on I27: Click in boxes to use drop-down lists
Note on I47: Click in boxes to use drop-down lists
On-site
Off-site
On-site
Off-site

## Tech QC Eval-GE Pristina
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(General Electric Pristina, continued)
Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.
|  | FREQUENCY | PASS/FAIL
1. | Acquisition Station Monitor Cleaning | Daily, Monthly (or as required)- GE Pristina 5845677-8EN, Rev. 2
2. | Viewbox and Viewing Conditions (NA if no hardcopy interpreted or compared) (NA GE Pristina 5845677-8EN, Rev. 2) | Weekly
3. | Image Uniformity and Bad Pixels (NA GE Pristina 5845677-8EN, Rev. 2) | Weekly
4. | Flat Field Test and Detector Calibration (GE Pristina 5845677-8EN, Rev. 2) | Weekly
5. | IQST | Weekly
6. | Phantom Image Quality | Weekly
7. | AOP and SNR Check | Monthly
8. | Visual Checklist | Monthly
9. | Repeat Analysis Check | Quarterly
10. | Compression Force Test | Semi-annually
11. | Reconstructed Image Uniformity Test (DBT) (NA GE Pristina 5845677-8EN, Rev. 2) | Weekly
12. | Flat Field Test and Detector Calibration (DBT) (GE Pristina 5845677-8EN, Rev. 2) | Weekly
13. | Phantom Image Quality - 3D (DBT) | Weekly
14. | AOP 3D Check (DBT) (NA GE Pristina 5845677-8EN, Rev. 2) | Monthly
15. | Review Workstation QC-Overall (NA if only hardcopy read) | See FDA guidance
16. | Film Printer QC (if applicable) | Printer mfr recommendations
17. | Mobile Unit Quality Control (if applicable) | After every move
Medical Physicist's Recommendations for Quality Improvement
Important: 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. 2. Use the QC manual version provided by the manufacturer for the digital system surveyed. 3. If the RWS or printer is FDA-cleared for FFDM, their QC manual is considered to be "substantially the same" and may be followed. (Check with the manufacturers for the QC manual and clearance status of their products.) 4. If the RWS or printer is not cleared by the FDA for FFDM, follow the QC manual provided by the image receptor manufacturer. (Check with the image receptor manufacturer for their required tests.) 5. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day than the survey date, note the date above. 6. See the FDA-approved alternative standard for GE FFDM regarding corrective action periods when components fail QC. However, if these tests are performed as part of a Mammography Equipment Evaluation (e.g., for a new system), corrective action must be taken before mammographic images are acquired.
Note on I10: Click in boxes to use drop-down lists

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5
Model
Pristina