## QC Test Summary-FDR
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – Fuji (FDR)
Site Name | [7 empty] | Report Date
Address | [7 empty] | Survey Date
Medical Physicist's Name | [7 empty] | Signature
X-Ray Unit Manufacturer |  |  | Fuji |  |  |  |  | Model
Date of Installation | [7 empty] | Room ID
QC Manual Version # | [6 empty] | (use version applicable to unit tested; contact mfr if questions)
Accessory Equipment |  |  | Manufacturer |  | Model |  | Location |  | QC Manual Version #
Review Workstation*
Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System www.fda.gov/CDRH/MAMMOGRAPHY/robohelp/START.HTM.
Survey Type |  | Mammo Eqpt Evaluation (MEE) of new unit (include MQSA Rqmts for Mammo Eqpt checklist) | [7 empty] | Annual Survey
Features |  | 2D |  | Digital Breast Tomosynthesis (DBT)
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
PASS/FAIL
1. 1Shot Phantom - Missed Tissue on Chest Wall Edge
| MEE Only Tests: |  |  | PASS/FAIL
| CNR |  |  |  |  | 1 Shot Phantom Sensitivity Constancy
| Geometric Distortion |  |  |  |  | System Artifact Evaluation
| Uniformity |  |  |  |  | Low Contrast Detectability (LCD)
| Dynamic Range |  |  |  |  | Linearity/Beam Quality Constancy
Spatial Resolution (SR)
2. ACR Phantom |  |  |  |  | Fibers | Specks | Masses
Phantom IQ Test on AWS
Phantom IQ Test on Printer
3. Image Basic Test
4. Compression Device Confirmation
5. Viewbox Maintenance
6. Monitor Quality Control (Secondary/AWS)
7. Additive Lag Effects
8. Multiplicative Lag Effects (Ghost)
9. Visual and Functional Test
10. Spatial Resolution (Magnification)
11. kVp Accuracy and Reproducibility
12. Half Value Layer (HVL)
13. Collimation Assessment
14. Radiation Output
15. AEC Reproducibility
16. CNR (all modes required for model)
17. AGD (all modes required for model)
18. AGD-ACR Phantom (≤3 mGy (300 mrad))
19. i-AEC System Performance Test (Aspire HD - NA)
20. X-ray Field at Chest Wall Edges (DBT)
21. Missed Tissue at Chest Wall Side (DBT)
22. In-Plane Resolution (DBT)
23. AEC Performance (DBT)
24. AGD (DBT)
25. Short Term Reproducibility (DBT)
26. Z-Resolution (DBT)
27. Review Workstation QC-Overall
28. Film Printer QC-Overall (if applicable)
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on M21: Click in boxes to use drop-down lists
Note on I29: Click in boxes to use drop-down lists
On-site
Off-site
On-site
Off-site

## Tech QC Eval-Fuji FDR
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(Fuji (FDR), continued)
Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.
|  | FREQUENCY | PASS/FAIL
1. | 1Shot Phantom | Weekly
2. | Good Practice (cleaning) | Weekly
3. | ACR Phantom | Weekly
4. | Monitor Quality Control (Secondary/AWS) | Weekly
5. | Repeat Analysis | Quarterly
6. | Compression Device Confirmation | Semi-annually
7. | ACR MAP Phantom (DBT) | Weekly
8. | Homogeneity (DBT) | Weekly
9. | Repeat Analysis (DBT) | Quarterly
10. | Review Workstation QC-Overall | See FDA guidance
11. | Film Printer QC (if applicable) | See FDA guidance
12. | Mobile Unit Quality Control (if applicable) | After every move
Medical Physicist's Recommendations for Quality Improvement
Important: 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. 2. Use the QC manual version provided by the manufacturer for the digital system surveyed. 3. If the RWS or printer is FDA-cleared for FFDM, their QC manual is considered to be "substantially the same" and may be followed. (Check with the manufacturers for the QC manual and clearance status of their products.) 4. If the RWS or printer is not cleared by the FDA for FFDM, follow the QC manual provided by the image receptor manufacturer. (Check with the image receptor manufacturer for their required tests.) 5. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day than the survey date, note the date above.
Note on I10: Click in boxes to use drop-down lists

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5
Model
Aspire
Aspire Cristalle
Cristalle