## QC Test Summary-GE
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital – General Electric (except Pristina)
Site Name | [7 empty] | Report Date
Address | [7 empty] | Survey Date
Medical Physicist's Name | [7 empty] | Signature
X-Ray Unit Manufacturer |  |  | General Electric |  |  |  |  | Model
Date of Installation | [7 empty] | Room ID
QC Manual Version # | [5 empty] | (use version applicable to unit tested; contact mfr if questions)
Accessory Equipment |  |  | Manufacturer |  | Model |  | Location |  | QC Manual Version #
Review Workstation*
Film Printer*
*FDA recommends that only monitors and printers specifically cleared for FFDM use by FDA’s Office of Device Evaluation (ODE) be used. See FDA's Policy Guidance Help System (www.accessdata.fda.gov/cdrh_docs/presentations/pghs/Polic_Guidance_Help_System.htm).
Survey Type |  | Mammo Eqpt Evaluation (MEE) of new unit (include MQSA Rqmts for Mammo Eqpt checklist) | [8 empty] | Annual Survey
Features |  | 2D |  | Digital Breast Tomosynthesis (DBT)
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
1. Phantom Image Quality |  |  |  |  | Fibers | Specks | Masses |  |  |  | PASS/FAIL
Phantom IQ Test on AWS
Phantom IQ Test on Printer (if applicable)
2. CNR Measurement (NA for DS or Essential/Ess Care if Sub-System MTF done)
| CNR |  | (required for new unit MEE and Annual Surveys)
Change in CNR ≤0.2 (NA for MEE)
3. MTF Measurement (NA for 2000D, DS or Essential/Ess Care if Sub-System MTF done)
4. AOP Mode and SNR Check
5. Collimation Assessment
6. Evaluation of Focal Spot Performance (NA for 2000D, DS or Essential/Ess Care if Sub-System MTF done)
7. Sub-System MTF (NA for 2000D if MTF and Focal Spot Performance done;
NA for DS or Essential/Ess Care if CNR, MTF and Focal Spot Performance done)
8. Breast Entrance Exposure, Average Glandular Dose and Reproducibility
| Average glandular dose for average breast is ≤3 mGy (300 mrad) | [8 empty] | mrad
Exposure reproducibility (CV) for air kerma (R) and mAs is ≤0.05
9. Artifact Evaluation and Flat Field Uniformity
10. Test for Flexible Paddle Deflection in Compression (if applicable)
11. kVp Accuracy and Reproducibility
12. Beam Quality Assessment (Half-Value Layer Measurement)
13. Radiation Output
14. Mammographic Unit Assembly Evaluation
Meets requirements for motion of tube-image receptor assembly
Meets requirements for compression paddle decompression
15. Phantom Image Quality (SenoClaire) |  |  |  |  | Fibers | Specks | Masses
Phantom IQ Test with MTD
Phantom IQ 3D Test
16. CNR and MTF Measurement with MTD (SenoClaire)
17. Flat Field 3D Test (SenoClaire DBT)
18. Grid Texture Test (SenoClaire)
19. AOP 3D Check (SenoClaire DBT)
20. Compression Paddle Border to Chest Wall Alignment with MTD (SenoClaire)
21. Breast Entrance Exposure and Average Glandular Dose (SenoClaire)
| MTD 2D Average glandular dose for average breast is ≤3 mGy (300 mrad) | [8 empty] | mrad
| 3D Average glandular dose for average breast is ≤3 mGy (300 mrad) | [8 empty] | mrad
22. Artifact Evaluation and Flat Field Uniformity with MTD (SenoClaire)
23. Volume Coverage (SenoClaire DBT)
24. Review Workstation (RWS) Tests (for all RWS, even if located offsite; NA if only hardcopy read)
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on I22: Click in boxes to use drop-down lists
Note on I45: Click in boxes to use drop-down lists
On-site
Off-site
On-site
Off-site

## Tech QC Eval-GE
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(General Electric - except Pristina, continued)
Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.
|  | FREQUENCY | PASS/FAIL
1. | Monitor Cleaning | Daily
2. | Flat Field Test | Weekly
3. | Phantom Image Quality | Weekly
4. | Viewbox and Viewing Conditions (NA if no hardcopy interpreted or compared) | Weekly
5. | CNR | Weekly
6. | MTF Measurement | Weekly (2000D-Monthly)
7. | AOP Mode and SNR Check | Monthly
8. | Visual Checklist | Monthly
9. | Repeat Analysis Check | Quarterly
10. | Compression Force Test | Semi-annually
11. | Phantom Image Quality - 2D (SenoClaire) | Weekly
12. | CNR and MTF Measurement with MTD (SenoClaire) | Weekly
13. | Flat Field 3D Test (SenoClaire DBT) | Weekly
14. | Phantom Image Quality - 3D (SenoClaire DBT) | Weekly
15. | Grid Texture Test (SenoClaire) | Monthly
16. | AOP 2D and SNR Check with MTD (SenoClaire) | Monthly
17. | AOP 3D Check (SenoClaire DBT) | Monthly
18. | Visual Checklist (SenoClaire) | Monthly |  | `
19. | Compression Force (SenoClaire) | Semi-annually
20. | Review Workstation QC-Overall (NA if only hardcopy read) | See FDA guidance
21. | Film Printer QC (if applicable) | Printer mfr recommendations
22. | Mobile Unit Quality Control (if applicable) | After every move
23. | Darkroom Cleanliness (if applicable) | Daily
24. | Processor QC (if applicable) | Daily
25. | Analysis of Fixer Retention (if applicable) | Quarterly
26. | Darkroom Fog (if applicable) | Semi-annually
Medical Physicist's Recommendations for Quality Improvement
Important: 1. The facility's "quality assurance program shall be substantially the same as the quality assurance program recommended by the image receptor [digital detector] manufacturer." This is required by the FDA. 2. Use the QC manual version provided by the manufacturer for the digital system surveyed. 3. If the RWS or printer is FDA-cleared for FFDM, their QC manual is considered to be "substantially the same" and may be followed. (Check with the manufacturers for the QC manual and clearance status of their products.) 4. If the RWS or printer is not cleared by the FDA for FFDM, follow the QC manual provided by the image receptor manufacturer. (Check with the image receptor manufacturer for their required tests.) 5. All tests must be evaluated for the facility's on and off-site equipment. If the evaluation was done on a different day than the survey date, note the date above. 6. See the FDA-approved alternative standard for GE FFDM regarding corrective action periods when components fail QC. However, if these tests are performed as part of a Mammography Equipment Evaluation (e.g., for a new system), corrective action must be taken before mammographic images are acquired.
Note on I10: Click in boxes to use drop-down lists

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5
Model
2000D
DS
Essential
Essential Care
Essential + SenoClaire