## QC Test Summary - S-F
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Screen-Film Systems
Site Name |  |  | Report Date
Address |  |  | Survey Date
Medical Physicist's Name |  |  | Signature
X-Ray Unit Manufacturer |  |  | Model
Date of Installation |  |  | Room ID
Film (mfr & type) |  |  | Screen (mfr & type)
Survey Type |  | Mammo Eqpt Evaluation of new unit (include MQSA Rqmts for Mammo Eqpt checklist) |  |  |  | Annual Survey
Medical Physicist's QC Tests
("Pass" means all components of the test passes; indicate "Fail" if any component fails. Tests must be done for both on and off-site equipment.)
PASS/FAIL
[5 empty] | ACR Guides | MQSA Regs
1. Mammographic Unit Assembly Evaluation
2. Collimation Assessment
Deviation between X-ray field and light field ≤2% of SID
X-ray field does not extend beyond any side of the IR by more than 2% of SID
Chest wall edge of compression paddle doesn't extend beyond IR by more than 1% of SID
3. Evaluation of Focal Spot Performance
Measured performance within acceptable limits for large focal spot
Measured performance within acceptable limits for small focal spot
4. Automatic Exposure Control (AEC) System Performance
Exposure reproducibility is within acceptable limits
AEC compensation for kVp, breast thickness and image mode is adequate
AEC density control function is adequate
5. Uniformity of Screen Speed
Optical density range is ≤0.30
6. Artifact Evaluation
Artifacts were not apparent or not significant
7. Phantom Image Quality Evaluation
4 largest fibers, 3 largest speck groups and 3 largest masses are visible
| Phantom image quality scores: |  | Fibers
Specks
Masses
8. kVp Accuracy and Reproducibility
Measured average kVp within +5% of indicated kVp
kVp coefficient of variation <0.02
9. Beam Quality (Half-Value Layer) Assessment
Half-value layer is within acceptable lower and upper limits at all kVp values tested
10. Breast Entrance Exposure, Average Glandular Dose and Radiation Output Rate
Average glandular dose for average breast is below 3 mGy (300 mrad)
| Average glandular dose to a 4.2-cm-thick breast on your unit is |  |  | mrad
| Radiation output rate is >800 mR/sec (7.0 mGy/sec) at 28 kVp with Mo/Mo |  |  | mR/sec
11. Viewbox Luminance and Room Illuminance
Mammographic viewbox is capable of a luminance of at least 3000 nit
Room illuminance (viewbox surface & seen by observer) is 50 lux or less
Important: If either test #7 (Phantom Image Quality Evaluation) or the Average Glandular Dose component of test #10 fail FDA's MQSA regulations, corrective action must be taken before any further examinations are performed. Corrective action must be taken within 30 days of the test date for all other MQSA failures.
*** YOUR MEDICAL PHYSICIST MUST SUMMARIZE HIS/HER RESULTS ON THIS FORM ***
Note on K17: Click in boxes to use drop-down lists
Note on H35: Click in boxes to use drop-down lists

## Tech QC Eval - S-F
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
(Screen-Film systems, continued)
Evaluation of Technologist QC Program
New units: Medical physicists must review the technologist QC within 45 days of installation and complete this section. The facility is required to submit the entire Mammography Equipment Evaluation report (including this form) along with their testing materials for accreditation.
Existing units: Medical physicists must complete this section as part of the unit's annual survey.
Relocating units: This section is not required if the medical physicist does not conduct a complete annual survey after relocation.
PASS/FAIL
|  | FREQUENCY | ACR Guides | MQSA Regs
1. | Darkroom cleanliness | Daily
2. | Processor QC (performed, records maintained, action as needed) | Daily
3. | Screen cleaning | Weekly
4. | Mammo phantom imaging (performed, records maintained, action taken as needed) | Weekly
5. | Darkroom fog | Semi-annually
6. | Film-screen contact test | Semi-annually
7. | Compression pressure monitored | Semi-annually
8. | Repeat analysis (performed, records maintained, reviewed by radiologist) | Quarterly
9. | Viewboxes and viewing conditions | Weekly
10. | Analysis of fixer retention | Quarterly
11. | Visual checklist | Monthly
Medical Physicist's Recommendations for Quality Improvement
Note on J11: Click in boxes to use drop-down lists

## dropdowns
Pass/Fail
Pass
Fail
Pass/Fail/NA
Pass
Fail
NA
Phantom Scores
6
5.5
5
4.5
4
3.5
3
2.5
2
1.5
1
0.5